## Slide 1: Sodium-borate (SB) vs. Tris-Borate-EDTA (TBE) for gel electrophoresis of DNA
By using SB for both gel and running buffer, we can:
1) Run gels faster & at higher-throughputs
2) Get better quality gels
-- Sharper bands
-- Fully separated bands in 100-bp DNA ladder
-- Better sample band separation & size estimates
-- Harder to overload with DNA
3) Save $ (pricing as of 21 April 2020)
-- TBE (1/2X) = $1.10/L (from 10X Liquid Concentrate 4L, VWR)
-- SB (1X) = $0.20/L (from powders/pellets, Sigma-Aldrich)

## Slide 2: Presentation overview: 30 slides
>> PCR product: 543 bp Alu (Owl monkey).
-- Compare TBE (Tris-Borate-EDTA) & SB (sodium borate) at various voltages & running times
-- show whole gels;
-- compare leaching effects of EtBr (ethidium bromide);
-- compare DNA overloading potential of TBE vs. SB;
-- focus on middle 2 rows in gels;
-- Remelting of SB gels.
-- SB gels at excessive voltage.
-- Recap TBE vs. best SB parameters.
-- High-throughput gels.
-- Summary slides.

## Slide 3: TBE (new gel): 45 min @ 200 V (standard run conditions)
Gel runs
-- Band cupping & overloading
-- Streaking and smearing of DNA in lanes
-- Poor separation at high MW
5 ul
1.5 ul
EtBr
leaching
200 bp
300 bp

## Slide 4: TBE (new gel): 15 min @ 600 V
Gel runs
-- Similar results, but somewhat more pronounced @ 600 V
5 ul
EtBr
leaching
200 bp
300 bp

## Slide 5: Sodium-borate: 15 min @ 600 V
Gel runs
-- Bands are straight
-- Minimal streaking
-- Good separation even at high MW
6 ul
2 ul
6 ul
12 ul
EtBr leaching

## Slide 6: Sodium-borate (used 1X): 10 min @ 600 V (“used” = re-melted gel, no extra EtBr added)
Gel runs
-- Similar results to 600 V for 15 min, except:
-- EtBr concentration ok... due to reduced run time and migration of EtBR from used buffer into gel front.
-- High MW bands not as well separated
6 ul
2 ul
EtBr
leaching
100 bp
200 bp

## Slide 7: Sodium-borate (used 1X): 15 min @ 500 V
Gel runs
-- Similar results to 600 V for 10 min, except here the High MW bands are well separated
6 ul
2 ul
EtBr
leaching
100 bp
200 bp

## Slide 8: Ethidium Bromide leaching
1) Ethidium bromide migrates in the opposite direction of the DNA; thus, the “front” of the gel can become devoid of ethidium bromide... causing DNA bands to ‘vanish’.
2) Primarily affects leading row of DNA samples
3) To Reduce effect when taking photos of leached gels:
a) move gel section away from edge of UV illuminator
b) crop photo, click on “transform”, and
adjust “gamma” slide

## Slide 9: Comparison of TBE & SB (time & voltage)
45 min
200V
15 min
600V
15 min
600V
10 min
600V
15 min
500V
TBE
SB
200 bp
200 bp
200 bp
300 bp
100 bp
100 bp
Bands are weak due to ‘leading edge’ effect.
Bands are stronger; less time for migration.
Bands are stronger; more time, but lower voltage.

## Slide 10: Overloading potential of TBE vs. SB gels
Sodium-borate (new): 15 min @ 600 V
5
1.5
6
12
ul DNA
ul DNA
TBE (new): 45 min @ 200V
Note that the SB gel gave much cleaner results, with crisper DNA bands... even with more DNA loaded in each lane.

## Slide 11: TBE vs. SB: Middle 2 rows of gels
TBE (new)
45 min
200V
SB (new)
15 min
600V
SB (used 1X)
10 min
600V
SB (used 1X)
15 min
500V
Again, SB gels gave much cleaner results, with crisper DNA bands... even with more DNA loaded in each lane.

## Slide 12: Remelting of SB gels: full gel
Sodium-borate (remelted 3 times): 15 min @ 600V – no issues

## Slide 13: Closer look at Remelting SB gels: 1st half gel
Sodium-borate (remelted 3 times): 15 min @ 600V – no issues

## Slide 14: Closer Look at Remelting SB gels: Leading edge of gel
Sodium-borate (remelted 3 times): 15 min @ 600 V – no issues
Note: EtBr leached into tray buffer from prior 3 gels apparently
“restained” the leading row of samples
100 bp

## Slide 15: High Voltage & SB Gels
Gel runs
T- 0 min: 164 mA; 98 W; 25oC
T-10 min: 260 mA; 156 W; 28oC (mA increased by <100; temperature increased 3oC)
Sodium-borate (used 1X): 10 min @ 600 V; minimal increase in mA & oC over run

## Slide 16: Excessive Voltage & SB Gels
750 V (10 min)
900 V (10 min)
– ↑’d 200 mA
– Temp ↑’d 12oC
>> Gel did not melt, but mA and temperature increases were excessive.
– 810 V by 5 min
– 700 V by 10 min
– ↑’d 200 mA
– Temp ↑’d 14oC
>> Parts of gel melted (see wavy bands in circle) and the change in gel consistency even affected the run voltage.
Gel runs

## Slide 17: Excessive voltage & EtBr leaching in SB gels
900 initial-V
750V
10 min
Moderately excessive voltage (750 V) caused rapid leaching of EtBr at gel front,
but just enough dye remained to still weakly stain the DNA bands.
Extreme excessive voltage (900 V) caused complete leaching of EtBr at gel front,
such that the DNA bands (molecular ladders and samples) are not visibly stained.

## Slide 18: Recap
-- Next 2 slides show results for
standard TBE vs. best SB parameters
-- In SB, DNA bands are still much sharper (more compact), and lanes show much less smearing.
Note: Remember to make BOTH the gel
    and the running buffer from
1X SB media.

## Slide 19: Gels for TBE vs. SB under each buffer’s “Best” run conditions
TBE (new)
45 min
200V
Sodium-borate (used 1X)
15 min
500V

## Slide 20: Full 25-cm SB gel: 4 rows; 26 wells/row
Sodium-borate (used 1X)
15 min
500V
Gel runs

## Slide 21: High-throughput with SB gels 10 combs in 25-cm gel (vs. 4)
Sodium-borate (new): 7 min @ 600 V
-- With TBE, no more than 4 combs could be used in a 25-cm gel; thus, each Population plate required 1 full gel.
-- Note: Gel was cut in ½ lengthwise for display purposes.
Gel runs
Pop Plate #2
Population Plate #1
½ Pop Plate #3
300 bp
500 bp
300 bp
500 bp

## Slide 22: Close-up view of High Throughput with SB gels
Sodium-borate (new)
7 min @ 600V
10 combs in 25-cm gel (vs. 4)
Gel runs
Pop Plate #2
Pop Plate #1
300 bp
500 bp
300 bp
500 bp
Note:
Band resolution would be even
better if run @ 300-400 V for ~15’.

## Slide 23: Sodium-borate (SB) Limitations
1) No advantages to 1/2X SB (unlike TBE)
-- runs significantly more slowly
-- gels don’t look as good as with 1X SB
-- This is because SB was specifically designed for running DNA. By contrast, 1X TBE was the recipe used for protein electrophoresis and turned out to be 2X stronger than needed for DNA electrophoresis.
2) Voltage for 25 cm gels:
-- 500 V (20 V/cm gel): quality is excellent.
-- 600 V (24 V/cm gel): quality is still good.
-- EtBr leaches from leading row by 15 min @ 600V.
-- 700 V (28 V/cm gel): quality deteriorates.
-- 900 V (36 V/cm gel): gels melt.
3) Gels <25 cm long?:
-- not tested here.
-- Brody & Kern: ran @ ≤35 V/cm for 10 cm gels (1.2%).

## Slide 24: Cloning from SB gels?
Brody & Kern (Biotechniques 36:214-216)
“We have performed gel extraction of restriction-
digested plasmid DNA from SB gels and successfully
subcloned the fragment.”
‘Faster Better Media’ website FAQ’s:
Q: “Will SB interfere with my DNA (sub-cloning, w/kits, etc.)?
A: “NO. We have used commonly used commercially available
kits to extract DNA and clone with SB. Your DNA might even
be happier in a cooler medium such as SB.”
Genomics Core (LSU, Dept. Biol. Sci.):
   We have had the same results.

## Slide 25: Advantages of SB running media
By using SB, we can:
1) Run gels at least 3X faster than with TBE
-- 15 min @ 500 V for extensive band separation
2) High-throughput of 2.5 pop plates/gel
-- 7 min @ 600 V for filled/empty site recognition
3) Get better quality gels
-- Sharper bands
-- Fully separated bands in 100-bp DNA ladder
-- Better sample band separation & size estimates
-- Harder to overload with DNA
4) Save $ (pricing as of 21 April 2020)
-- TBE (1/2X) = $1.10/L (from 10X Liquid Concentrate 4L, VWR)
-- SB (1X) = $0.20/L (from powders/pellets, Sigma-Aldrich)

## Slide 26: Recommendations by Brody & Kern
1) Standard applications:
-- 10 mM sodium boric acid (SB)
2) Small DNA or RNA fragments (1 bp resolution):
-- 1 mM lithium boric acid
3) High resolution of large DNA fragments (>3 kb):
-- 5 mM lithium acetate
 BioTechniques 36:214-216 (February 2004). Sodium boric acid:
a Tris-free, cooler conductive medium for DNA electrophoresis.
 Analytical Biochemistry 333 (2004) 1–13. History and principles of
conductive media for standard DNA electrophoresis.
 http://www.fasterbettermedia.com/

## Slide 27: Sources for SB media
“Powder”: 1X @ $0.20/L
Make 20X solutions in Glass bottles from either:
A) 5 mM disodium borate decahydrate; or,
B) 10 mM sodium hydroxide pH’d with borate, as follows:
-> 8 g NaOH pellets ($70/500g, Sigma-Aldrich)
-> 36.76 g H3BO3 ($78/kg, Sigma-Aldrich)
-> Stir & heat in 800 ml Nanopure water;
cool, qs to 1000 ml (~ pH 8.5 @ RT).
-> 1 kg H3BO3 and 217 g NaOH are sufficient to make 544 L of 1X SB for ~$0.2/liter.
-> Solution must be heated to dissolve the boric acid.
-> If your usage is high, SB can be prepared as 40X solutions; however, precipitates will form within
a few weeks even in glass bottles; in that case, reheat to dissolve prior to making 1X dilution.
-> By contrast, 20X solutions do not precipitate in glass bottles stored at room temperature.
2) 20X SB™: Sodium boric acid conductive medium
from Faster Better Media. – 21Apr2020 prices, shipping is extra.
-- 1X @ $4.15/L: $83, 1 liter bottle.
-- 1X @ $3.05/L: $244, 4 liter box, spigot.
-- 1X @ $2.28/L: $455, 10 liter box, spigot.
-- 1X @ $2.04/L: $816, 20 liter box, spigot.

## Slide 28: Finis