Agilent RNA Pico 6000  (04 June 2025; Scott Herke)

See “Science Aid Center” on Core website for additional information as indicated below.

Cleaning the Electrodes before Running Assays

- Unless RNAse contamination is expected, skip to Step 5, with 5’ incubations.

If RNAse contamination is suspected, do Steps 2-4 before running any RNA Pico assays.

- Slowly fill RNAseZAP electrode cleaner with 340 µl of fresh RNaseZAP.

- Place electrode cleaner in Bioanalyzer, close lid, and incubate (5 min).  	___|

- Remove electrode cleaner, flick out RNAseZAP, and store cleaner chip.

>> Repeat Steps 5-7, ≥4 times:

- Slowly fill the RNA-WATER electrode cleaner with 340 µl of fresh RNase-free water.

- Place electrode cleaner in Bioanalyzer, close lid, & incubate.

  - Normal procedure: 	___   5 minute incubations.

  - RNAseZAP procedure:	___ 10 minute incubations.			___|___| ___| ___|

- Remove electrode cleaner, flick out water.

>> After the 3-4X repetitions, store cleaner chip.

- Leave lid open (5 min) to allow water to evaporate from electrodes.

Basic start-up procedures:

- Before working with the chip, clean work areas with 80% EtOH (to minimize dust which can contaminate the chip or reagents with environmental RNAse).

- Periodically, spritz gloves with RNAse-ZAP; rub gloves together until ‘dry’.

- Start software before you load the chip; verify that a connection is live.

- Adjust syringe clip:  Release clip lever and slide it up to the TOP position.

- Inspect white seal of chip priming station to ensure no old gel is plugging the hole; clean if needed.

- Allow all reagents to equilibrate to room temperature for 30 minutes before use.

- Protect dye and dye mixtures from light.

- Keep samples on ice.

- RNA ladder: remove heat-denatured ladder aliquot from -70°C freezer, and thaw on ice... do NOT allow ladder to warm extensively.

- Heat-denature (70°C, 2’) RNA samples & snap-chill, unless samples might have ‘hidden breaks’ (see “Hidden Breaks (28S RNA peak)” ... in which case, it might be best to skip the denaturation step.

- Always vortex dye concentrate (10 s) before preparing the gel- dye mix and spin down afterwards.

- Set the timer to 30 s (not 1 minute, as for the DNA-HS assay) and vortexer to 2000.

LADDER PREPARATION:  Area under ladder is used to quantitate samples... dilute correctly!

To avoid RNase contamination and repetitive freeze/thaw cycles, RNA ladder must be aliquoted in RNase-free vials.

➔ Some plastics can bind RNA. This can affect RNA ladder concentration which affects the ladder identification and the quantitation of the sample. Use Eppendorf Safe-Lock PCR clean or Eppendorf DNA LoBind 0.5ml Microcentrifuge tubes for RNA ladder aliquoting.

After reagent kit arrival, prepare a work area by cleaning with 80% EtOH and then RNAseZAP, and then...

- Spin ladder down and transfer to 0.2-ml Axygen DNA LoBind PCR tube.

- Heat denature the ladder for 2 min (70°C) in thermal cycler.

- Immediately cool the vial on ice.

- Following the heat-denaturation, add 90 µL of RNase- free water and mix thoroughly.

- Prepare aliquots in 0.5 mL RNase- free vials (30) with 3.2 ul (each) for typical daily use.

- Store aliquots at -70°C.

- Before use, thaw ladder aliquots and keep on ice (avoid extensive warming upon thawing process).

Agilent RNA Pico 6000 – Page 2 of 2

Preparing gel:  (Note:  If the previously prepared gel is >1 month old, repeat the filtering step.)

- Add 550 ul gel to spin filter and centrifuge (10’, @ 1200-1800 rcf).

- Aliquot 65 µl filtered gel into a 0.5 ml RNase- free microcentrifuge tube.

- Store remainder in the 1.5-ml tube at 4°C.

Adding dye concentrate to gel:

- Vortex RNA 6000 Pico dye concentrate (blue) for 10 seconds and spin down.

- Add 1 µl of RNA 6000 Pico dye concentrate (blue) to a 65 µl aliquot of filtered gel.

  - Return dye concentrate to 4 °C (dark).

- Cap tube, vortex thoroughly and visually inspect proper mixing of gel and dye.

- Spin tube (10’ @ RT, 13000 rcf).

- Use prepared gel-dye mix within 1 day (discard if not used).

Preparing samples and Ladder:

- To minimize 2o-structure, heat-denature (70°C, 2’) & snap-chill samples before loading on chip.

(Note: If samples might have ‘hidden breaks’, it might be best to skip the denaturation step.)

- Thaw ladder aliquot and keep on ice (avoid extensive warming upon thawing process).

- Pipette 1 µl of the diluted RNA 6000 Pico ladder into a 0.2-ml tube.

- Pipette 1 µl of each sample (or water) into 11 tubes (0.2-ml).

- Centrifuge 12 tubes; add 5 ul RNA 6000 Pico marker (green) to each tube; briefly vortex; touch-spin.

Loading the chip:

- Adjust syringe clip:  Release clip lever and slide it up to the TOP position.

- Inspect white seal of chip priming station to ensure no old gel is plugging the hole.

- Remove a new RNA chip from bag and place it on the chip priming station.

- Pipette 10.0 µl of gel-dye mix at bottom of ‘loading’ well.

  - When pipetting gel-dye mix, do NOT draw up particles that may sit at the bottom of the gel-dye mix vial.

  - To prevent forming a large air bubble under the gel-dye mix, insert tip of pipette to bottom of chip well when dispensing. (Placing pipette at edge of well may lead to poor results.)

- Set timer (30 s), make sure that plunger is positioned at 1 ml and then close chip priming station.

- Press syringe plunger down until it is held by the clip.

- Wait for exactly 30 seconds and then release plunger with the clip release mechanism.

- Visually inspect that the plunger moves back at least to the 0.3 ml mark almost immediately.

- Wait for 8-10 seconds, then slowly pull back plunger to 1 ml position.

- Open chip priming station – observe white sealing disc with magnifying glass; blot disc with Kimwipe if any gel is detected (prevents clogging of the chip priming station and subsequent chip failures).

- Pipette 10.0 µl of gel-dye mix into the other two “G” wells.

- Pipette 10 µl of RNA 6000 Pico conditioning solution (white) into the “CS” well.

- Pipette samples/ladder into the 12 wells.  (Note: Do not leave any wells empty or the chip will not run properly.

Final chip preparation

- Place chip in IKA vortex mixer (ensuring it is flat and securely held); vortex (60 s, 2000 rpm).

- Centrifuge chip (30 s, 100 rcf) in a Plate Centrifuge (using a cardboard or 3D-printed ‘chip holder’ as shown at “Loading Agilent Bioanalyzer chips... Genomics Core style!”.

- Place chip in Bioanalyzer, close lid, and start the run (within 5’ of chip loading).

- After the run is finished, remove chip and clean [2X] electrodes (steps 5-7 of Cleaning procedure).

Do not leave chip on Bioanalyzer for >1 hour, or electrodes may need vigorous cleaning to remove dried contaminates.